The protocol between Power supply and PC
Power supply can be connected to RS-232 interface by the DB9 plug on the back panel through the level switching circuit. The following instructions can help you to understand how to control the output of power supply by PC.
A. RS-232 Communication Set
You can set the baud rate and address of the power supply by the MENU on the LCD panel.

1． Address: (0—31)

2． Baud rate: 9600 (4800, 9600, 19200, 38400)
3． Data bit: 8

4． Stop bit: 1

5． Verify: None

[image: image1.png]



B. DB9 Serial interface
The output of DB9 interface on the back panel is TTL, and you need to level switching it by the auxiliary 3311 to transfer and connect it to the PC Com port.

Power          Level Converter              3110         PC

[image: image2.wmf]1

2

3

4

5

6

7

8

9

VCC

RXD

TXD

NC

GND

NC

NC

NC

NC

8

6

1

2

7

9

4

5

3

RTS

RXD

NC

GND

DTR

NC

NC

TXD

VCC


C. Frame format
  The length of the frame is 26 bytes (compatible with FAB). The form is as the following:

	AAH
	Address
	Command
	Byte 4 to 25 are relevant information
	Check 


Directions: 

1． AAH occupies a byte
2． Address (0~31) occupies a byte
3． Command occupies a byte. And the contents are as the following: 
a． 80H------Set the upper limit of current, power and the voltage level.
b． 81H------Read the current, voltage, power and the current status. The statuses include ON/OFF, over current and overpower status of the power supply. 
c． 82H------Control the ON/OFF of the power supply
d． 83H------Set the protection of power supply
e． 84H------Read the protection status of power supply
f． 85H------Demarcate the power command
g． 86H------Return the actual output voltage to power supply 
h． 87H------Demarcate the current command
i． 88H------Return the actual output voltage to power supply
j． 89H------Set the demarcating information of power supply
k． 8AH------Read the demarcating information
l． 8BH------Set the serial number of power supply
m． 8CH-----Read the serial number, product model and software version of power supply
n． 12H-------Check

4． If you want to control the output of the power supply by PC, you have to set the power supply as PC controlled status. And the command is 82H. If you want to demarcate the output of the power supply and set the demarcating information and serial number of the power supply, you have to set the protection status as OFF first.
5． The 26th byte is check. It’s the sum of the previous twenty-five bytes.
D. The instructions of the Command list
   1．Set Max current, Max power and voltage levgel (80H)

	Byte 1
	AAH

	Byte 2
	Addresses (0—31) 

	Byte 3
	Command (80H)

	Byte 4
	Low byte of the Max current

	Byte 5
	High byte of the Max current

	Byte 6
	Low byte of the low word of the Max voltage

	Byte 7
	High byte of the low word of the Max voltage

	Byte 8
	Low byte of the high word of the Max voltage

	Byte 9
	High byte of the high word of the Max voltage

	Byte 10
	Low byte of the Max power

	Byte 11
	High byte of the Max power

	Byte 12
	Low byte of the low word of the voltage set

	Byte 13
	High byte of the low word of the voltage set

	Byte 14
	Low byte of the high word of the voltage set

	Byte 15
	High byte of the high word of the voltage set

	Byte 16
	New address of the power supply

	Byte 17~25
	Keep for the system

	Byte 26
	Check


The Current, Voltage and Power are all expressed in two bytes, with low byte in the front and the high byte behind. 

For example: The current value 3589H is showed as
	89H
	35H


The range of current set：0-3000mA  

The range of voltage set：0-36000mV 

The range of power set: 0-108 (Enlarge 100 times)

2．Read the current, voltage, power value and the status of the power supply (81H)

	Byte 1
	AAH

	Byte 2
	Address (0—31)

	Byte 3
	Command（81H）

	Byte 4
	Low byte of current

	Byte 5
	High byte of current

	Byte 6
	Low byte of the low word of the voltage

	Byte 7
	High byte of the low word of the voltage

	Byte 8
	Low byte of the high word of the voltage

	Byte 9
	High byte of the high word of the voltage

	Byte 10
	Low byte of power

	Byte 11
	High byte of power

	Byte 12
	Low byte of Max current

	Byte 13
	High byte of Max current

	Byte 14
	Low byte of the low word of the Max voltage 

	Byte 15
	High byte of the low word of the Max voltage 

	Byte 16
	Low byte of the High word of the Max voltage 

	Byte 17
	High byte of the High word of the Max voltage

	Byte 18
	Low byte of Max power 

	Byte 19
	High byte of Max power 

	Byte 20
	Low byte of the low word of the voltage set

	Byte 21
	High byte of the low word of the voltage set

	Byte 22
	Low byte of the High word of the voltage set

	Byte 23
	High byte of the High word of the voltage set 

	Byte 24
	Power supply status 

	Byte 25
	Keep for the system 

	Byte 26
	Checking code 


The current, voltage and power are all expressed in two bytes, with low byte in the front and high byte behind
Power supply status occupies one byte, each bit is defined as: 

          From high to low 

	7
	6
	5
	4
	3
	2
	1
	0


  Bit 0：Power supply status，0 for OFF，1 for ON
Bit 1：Over current status，0 for Normal，1 for abnormal
Bit 2：Over power status，0 for Normal，1 for abnormal
Bit 3：Operation status，0 for Keyboard，1 for PC

Notes: The frame format from Power supply to PC is as the above.

3．Control the ON/OFF of the power supply (82H)
	Byte 1 
	AAH

	Byte 2
	Address (0—31)

	Byte 3
	Command (82H)

	Byte 4
	The status of the power supply

	Byte 5~25
	Keep for system

	Byte 266
	Check6


The status of the power supply is expressed in one byte. Each bit is defined as the following: 
                 From high to low  

	7
	6
	5
	4
	3
	2
	1
	0


  Bit 0： The status of the power supply，0 for OFF，1 for ON

Bit 1： The power supply controlled by PC，0 for Self-controlling of the power itself，1 for PC controlling the power
Explanation：Only under the situation of PC controlling, you can set the parameters of the power.

4. Power supply automatically upper-transmits the max current, the max power and the voltage levels to the PC.

The frame is the same as frame of the setup of the max current, the max power and the voltage levels.
Examples: 

1. Set the parameters:

3000mA,36000mV,10800(108W) ,3000V

AA 00 80 B8 0B A0 8C 00 00 30 2A B8 0B 00 00 00 00 00 00 00 00 00 00 00 00 36
2. Read the parameters: 

AA 00 81 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 2B
3. Set control status:
A: PC control, output ON

AA 00 82 03 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 2F

B: PC control, output OFF

AA 00 82 02 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 2E

4.Self-controlled

AA 00 82 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 2C
_1108876944

